## Slide 1: Low-Power ECG-Based Processor for Predicting Ventricular Arrhythmia

## Slide 2: Abstract
This paper presents the design of a fully integrated electrocardiogram (ECG) signal processor (ESP) for the prediction of ventricular arrhythmia using a unique set of ECG features and a naive Bayes classifier. Real-time and adaptive techniques for the detection and the delineation of the P-QRS-T waves were investigated to extract the fiducial points. Those techniques are robust to any variations in the ECG signal with high sensitivity and precision. Two databases of the heart signal recordings from the MIT PhysioNet and the American Heart Association were used as a validation set to evaluate the performance of the processor. Based on application-specified integrated circuit (ASIC) simulation results, the overall classification accuracy was found to be 86% on the out-of-sample validation data with 3-s window size. The architecture of the proposed ESP was implemented using 65-nm CMOS process. It occupied 0.112-mm2 area and consumed 2.78-µW power at an operating frequency of 10 kHz and from an operating voltage of 1 V. It is worth mentioning that the proposed ESP is the first ASIC implementation of an ECG-based processor that is used for the prediction of ventricular arrhythmia up to 3 h before the onset.

## Slide 3: Existing System
The sequence of electrical activities are reflected into the ECG waveform, and therefore, temporal dependencies naturally exist in this waveform.
LSTM solves this issue by allowing to forget according to the actual dependencies which exist in the problem.
The first part, i.e., local data, is specific to the patient and is helpful in increasing the classification accuracy due to existing similarities among the heartbeats of every patient.
This method cannot be applied to RNNs because of the existing temporal dependencies in the model, i.e., the feedback loops in which carry previous information through time.

## Slide 4: Disadvantage of Existing System
Ventricular arrhythmia is an abnormal ECG rhythm and is responsible for 75%–85% of sudden deaths in persons with heart problems unless treated within seconds.
The main benefit of such approach is that the impact of a person’s motion and his daily activities is dramatically reduced.
The main disadvantage of the system is that it uses fixed search window with predefined size to locate S and T fiducial points, which is not suitable for real-time scenarios.
This computationally intensive algorithm is designed for a different problem namely classifying ECG signals into rhythms such as Sinus and Bigeminy.

## Slide 5: PROPOSED SYSTEM :
To extract the features automatically and increase the heartbeat classification accuracy, deep-learning based algorithms including deep convolutional neural networks and recurrent neural networks have recently been proposed.
The proposed algorithm employs RNNs because the ECG waveform is naturally fit to be processed by this type of neural network.
Our proposed algorithm is evaluated using the same ECG signals that were employed in the previous works that conform to this standard.
In contrast, our proposed method is designed from ground up to be lightweight, and hence, meets timing requirements for continuous execution on wearable devices with limited processing capacity.

## Slide 6: ADVANTAGE
These methods have exhibited advantages in the detection of ventricular arrhythmia, they have some shortcomings.
It was intended to investigate the performance of the system without introducing the strong biasing effect of a classifier.
The system is verified to operate for different clock frequencies, and we have reported the performance for operating frequency ranging from 10 kHz up to 4 MHz.
The performance of the technique is highly accurate and satisfactory for advanced ECG feature extraction highlights the capability of the system to detect and delineate the ECG signal in different real-life scenarios.

## Slide 7
TITLE | AUTHOR | DESCRIPTION
Ventricular arrhythmias: State of the art | J. W. Schleifer and K. Srivathsan, | The management of ventricular tachycardia and ventricular fibrillation in the cardiac intensive care unit can be complex.
Sudden cardiac death | D. P. Zipes and H. J. J. Wellens, | This article looks into the causes and methods of preventing SCD and at some of the Indian data.

## Slide 8
TITLE | AUTHOR | DESCRIPTION
Automatic classification of heartbeats using ECG morphology and heartbeat interval features | P. de Chazal, M. O’Dwyer, and R. B. Reilly, | A method for the automatic processing of the electrocardiogram (ECG) for the classification of heartbeats is presented.
Detecting ventricular fibrillation by time-delay methods | A. Amann, R. Tratnig, and K. Unterkofler, | A pivotal component in automated external defibrillators (AEDs) is the detection of ventricular fibrillation (VF) by means of appropriate detection algorithms.

## Slide 9: Block Diagram

## Slide 10: SOFTWARE REQUIREMENTS:
Operating System : Windows 10
Platform : DOT NET TECHNOLOGY
Front End : ASP.Net 4.0
Back End : SQL SERVER 2014

## Slide 11: HARDWARE REQUIREMENTS:
Keyboard
Mouse
Hard disk 500GB
Ram 4 Gb

## Slide 12: CONCLUSION
In this paper, a fully integrated digital ESP for the prediction of ventricular arrhythmia that combines a unique set of ECG features with naive Bayes was proposed. Real-time and adaptive techniques for the detection and delineation of the P-QRS-T waves were investigated and employed to extract the fiducial points. Furthermore, seven features that represent different intervals of the ECG signal were extracted and used as input to the naive Bayes to classify each heartbeat as normal or abnormal. The combination of these features has never been used in any previous detection or prediction system. The ESP was implemented using the state-of-the-art 65-nm technology, and based on the design constraints, it occupied an area of 0.112 mm2 and consumed a total power of 2.78 μW. Moreover, the proposed ESP achieved an outstanding capability of predicting the arrhythmia up to 3 h before the onset. A prediction accuracy of 86% was obtained on the out-of-sample validation data by tenfold cross validation with 3-s window size. The small area, low power, and high performance of the proposed ESP make it suitable for inclusion in system on chips targeting wearable mobile medical devices. The improvements that will be added to this design in the future could be summarized as follows. 1) Integrate multiple biomedical signals with the ECG, such as blood glucose, electroencephalograph, and electromyography. 2) Include a low-powered wireless transceiver module to transmit the biomedical signals. 3) Employ different filtering techniques to remove any type of noise that could be coupled with the ECG signal. 4) Employ a search-back mechanism in case of miss detection of any of the P-QRS-T waves. 5) Employ some power reduction techniques, such as clock and frequency scaling.

## Slide 13: REFERENCE
[1] J. W. Schleifer and K. Srivathsan, “Ventricular arrhythmias: State of the art,” Cardiol. Clin., vol. 31, no. 4, pp. 595–605, 2013.
[2] D. P. Zipes and H. J. J. Wellens, “Sudden cardiac death,” Circulation, vol. 98, no. 21, pp. 2334–2351, 1998.
[3] C. J. Garratt, Mechanisms and Management of Cardiac Arrhythmias. London, U.K.: BMJ Books, 2001.
[4] P. de Chazal, M. O’Dwyer, and R. B. Reilly, “Automatic classification of heartbeats using ECG morphology and heartbeat interval features,” IEEE Trans. Biomed. Eng., vol. 51, no. 7, pp. 1196–1206, Jul. 2004.
[5] A. Amann, R. Tratnig, and K. Unterkofler, “Detecting ventricular fibrillation by time-delay methods,” IEEE Trans. Biomed. Eng., vol. 54, no. 1, pp. 174–177, Jan. 2007.

## Slide 14
[6] O. Sayadi, M. B. Shamsollahi, and G. D. Clifford, “Robust detection of premature ventricular contractions using a wave-based Bayesian framework,” IEEE Trans. Biomed. Eng., vol. 57, no. 2, pp. 353–362, Feb. 2010.
[7] X.-S. Zhang, Y.-S. Zhu, N. V. Thakor, and Z.-Z. Wang, “Detecting ventricular tachycardia and fibrillation by complexity measure,” IEEE Trans. Biomed. Eng., vol. 46, no. 5, pp. 548–555, May 1999.
[8] J. Pardey, “Detection of ventricular fibrillation by sequential hypothesis testing of binary sequences,” in Proc. IEEE Comput. Cardiol., Sep./Oct. 2007, pp. 573–576.
[9] Q. Li, C. Rajagopalan, and G. D. Clifford, “Ventricular fibrillation and tachycardia classification using a machine learning approach,” vol. 61, no. 3, pp. 1607–1613, Jun. 2013.
[10] B.-Y. Shiu, S.-W. Wang, Y.-S. Chu, and T.-H. Tsai, “Low-power low-noise ECG acquisition system with dsp for heart disease identification,” in Proc. IEEE Biomed. Circuits Syst. Conf. (BioCAS), Oct./Nov. 2013, pp. 21–24.